Summary of Calibration Factors

C-130 - Basic NOMADSS Configuration

Guenther et al

Pre-Calibration – 05/14/2013

Param

Variable                    Name      C1        C2       C3

Dynamic Pressure-Fuselage L. (hPa)	QCF  	   1.318	  17.257  -.0032

Dynamic Pressure-Radome (hPa)	QCR 	  -1.113	  17.359	-.0035

Dynamic Pressure-Fuselage R. (hPa)	QCFR	  -0.17	  13.848   .0079

Total Temperature-Radome Left (oC)	TTRL	 -62.254	  23.552	 .0140

Total Temperature 1:Heated Fuse(oC)  TTHL1   -81.870    22.636   .3040

Total Temperature 2:Heated Fuse(oC)  TTHL2   -82.095    22.705   .2870

Remote Surface Temp. Bot. (oC) 	RSTB    -55.750    10.763  -.0164

Remote Surface Temp.-Bot. (oC) 	RSTB1   -46.998     8.646   .1228

Remote Surface Temp. Top. (oC)	RSTT    -49.385     9.014   .1123

Vertical Press. Diff. Radome(mb)	ADIFR	  -0.008     7.018	 .0022

Horizontal Press. Diff. Radome(mb) 	BDIFR	  -0.468     5.154   .0000

CN Flow Rate (slpm)	FCN	   0.007	   0.135   .1530

CN Inlet Pressure (hPa)	PCN	   3.110   108.470  -.0230

Cabin Pressure (hPa)	PCAB     -1.490   108.902  -.0466

Post Calibration – 07/29/2013

Param

Variable                    Name      C1        C2       C3__

Dynamic Pressure-Fuselage L. (hPa)	QCF#  	   1.264	  17.277   .0010

Dynamic Pressure-Radome (hPa)	QCR# 	  -1.127	  17.359	-.0035

Dynamic Pressure-Fuselage R. (hPa)	QCFR#	  -0.451    13.858  -.0026

Total Temperature-Radome Left (oC)	TTRL	 -62.234	  23.579	 .0160

Total Temperature 1:Heated Fuse(oC)  TTHR1   -82.059    22.620   .3057

Total Temperature 2:Heated Fuse(oC)  TTHR2   -82.167    22.611   .3013

Vertical Press. Diff. Radome(mb)	ADIFR	   0.107     6.994	 .0057

Horizontal Press. Diff. Radome(mb) 	BDIFR	   0.034     5.157   .0003

# Post Cal used in production processing

* Water Calibration

Calibration Maneuvers –

RAF uses sets of in-flight maneuvers specifically designed to help evaluate the performance of the 3-D wind measurement system on the aircraft.  Sets of maneuvers were flown on one test flight and one research flight: TF01 and RF11 respectively. RF11 also included an intercomparison interval with the NOAA P-3. In each case the variations observed fell within acceptable limits.

Calibration Summary

page 2

PMS - 1-D Probe Bin Sizes

The SPP100 (std), SPP200, CDP and SPP300 probes were used on the project.  The minimum particle size (microns) for each sampling bin are listed below:

UHSAS		 SPP200          CDP  		SPP300

bin#1 		void   um		 0.097 um		  2.25 um		0.266

bin#2		 	void		 	 0.105		  3.25		0.284

bin#3			void	 		 0.113		  4.25		0.32

bin#4			void			 0.121		  5.25		0.371

bin#5			void			 0.129		  6.25		0.423

bin#6			void			 0.145		  7.25		0.475

bin#7			void			 0.162		  8.25		0.528

bin#8			void			 0.182		  9.25		0.581

bin#9			void			 0.202		 10.25		0.635

bin#10		void			 0.222		 11.25		0.7

bin#11		void 		 0.242		 12.25		1.06

bin#12		0.084		 0.262		 13.25 		1.795

bin#13		0.087		 0.282		 14.25		2.43

bin#14		0.089		 0.302		 16.125		3.449

bin#15		0.092		 0.401		 18.0		4.112

bin#16		0.0945		 0.570	  	 19.875		4.792

bin#17		0.097		 0.656		 21.75		5.578

bin#18		0.100		 0.740	  	 23.625 		6.668

bin#19		0.103		 0.833		 25.5		7.728

bin#20		0.106		 0.917		 27.375		8.717

bin#21		0.109		 1.008		 29.25		9.665

bin#22		0.112		 1.148		 31.125	    10.651

bin#23		0.115		 1.319		 33.0	    11.636

bin#24		0.119		 1.479		 34.875	    12.623

bin#25		0.122		 1.636		 36.75	    13.614

bin#26		0.126		 1.796		 38.625	    14.647

bin#27		0.129		 1.955		 40.5	    15.595

bin#28		0.133		 2.184		 42.375	    16.683

bin#29		0.137		 2.413		 44.25	    17.776

bin#30		0.141		 2.661		 46.125	    18.876

bin#31		0.145		 2.991		 48.0	    20.0

.			  .

.              .

.              .

Bin#100		1.000