HIPPO-4 and HIPPO-5 project manager data QC report

Variable list

Attachments: hippo-4 variable list.txt; HIPPO-5 variable list.txt

Instrumentation

Pressure: 
Static pressure is available using two different systems: Research and Avionics.

Research static pressure is measured with a Paroscientific (MODEL 1000) with a stated accuracy of 0.01% of full scale. This measurement is output in the netCDF files as:

| PSF | static pressure as measured using the fuselage holes |
| --- | --- |
| PSX | same as PSF. Used to choose reference variable if more than one instrument provides measurement of the same parameter |
| PSFC | static pressure corrected for airflow effects (pcor) |
| PSXC | same as PSFC. Used to choose reference variable if more than one instrument provides measurement of the same parameter |

Use PSXC for the normal measure of pressure (e.g., in equation of state or hydrostatic equation).

Avionics static pressure is recorded from the GV avionics. This is slower than the Paroscientific measurement, but it has been corrected for airflow effects and it is certified for "Reduced vertical separation minimum" (RVSM) through the calculation of pressure altitude.  No documentation is available on how Gulfstream and Honeywell corrected this proprietary pressure measurement but it has passed very strict FAA certification requirements.

| PS_A | Avionics static pressure |
| --- | --- |

Temperature: 
Temperature was measured using five different sensors on the GV:

Dual element heated Harco and dual element heated Rosemount sensors were used to measure the total temperature. These sensors are characterized by a moderate response rate but are unaffected by icing. A fifth measurement of temperature (slow and with some delay) was provided by the GV avionics instrumentation.

The Harco measurements were logged using analog channels, subject to a three stage calibration (sensor head bath calibration, A/D board calibration, final engineering calibration) and are affected by a variable recovery factor. The recovery factor is a function of altitude and mach number, and RAF is currently fine tuning the correction algorithms for the recovery factor.  A/D corrections are analog board specific and are logged in the project calibration files and accompanying RAF project hardware documentation.

There are currently complications of A/D calibrations and instability of the thermal environment of the GV DSM and amplifiers used to collect temperature data. A comparison with multiple radiosondes led RAF to believe that the avionics temperature is the most accurate measurement at this time. ATX was set to AT_A for this project.

| TTHR1 | Total air temperature from the heated HARCO sensor # 1 right |
| --- | --- |
| TTHR2 | Total air temperature from the heated HARCO sensor # 2 right |
| TTFH1 | Total air temperature from the heated Rosemount sensor # 1 left |
| TTFH2 | Total air temperature from the heated Rosemount sensor # 2 left |
| TT_A | Total air temperature from the avionics system |
| ATFR | Ambient air temperature from the Rosemount system |
| ATHR1 | Ambient temperature from the heated HARCO sensor # 1 right |
| ATHR2 | Ambient temperature from the heated HARCO sensor # 2 right |
| ATFH1 | Ambient temperature from the heated Rosemount sensor # 1 left |
| ATFH2 | Ambient temperature from the heated Rosemount sensor # 2 left |
| AT_A | Ambient temperature from the avionics system |
| ATX | Ambient temperature reference. This is usually the same as AT_A but can be replaced by another sensor output if AT_A experiences a problem on a particular flight. In such case a flight-by-flight report will note the change. |

RAF recommends using ATX for the temperature in thermodynamic equations, etc.

Dewpoint temperature and vapor density:

Humidity was measured using two Buck Research 1011C cooled-mirror hygrometers that are normally used for measuring tropospheric humidity.  They have a sandwich of three Peltier elements to cool the mirror, and in comparison to earlier generations of cooled-mirror hygrometers, they have a much-improved capability to measure at low temperatures.  These sensors are assumed to measure dewpoint above 0°C and frostpoint below 0°C.  The instrument has a quoted accuracy of 0.1 °C over the -75 to +50 °C; however, based on examination of the measurements RAF is not comfortable with accuracies better 0.5 °C for dewpoint and 1 °C for frostpoint.  The cooled-mirror sensors have a slow response, in particular at lower temperatures, and this may give considerable differences between the measurements from the two units or when comparing with faster instruments.  Their cooling rates depend in part on the airflow through the sensor, and this may depend on the angle of the external stub relative to the airflow.  The angle may differ between the two sensors, and this may contribute to response-time differences between the sensors.  At very low temperatures the sensors may jump ("rail") to even lower temperatures; these data were removed from the dataset and replaced with a NAN.  The water vapor concentration derived from the chilled mirror sensors is used in other calculations (e.g., true airspeed). However, the impact of these out of bounds conditions on derived calculations that depend on humidity correction is very small at extremely low dew points, so omitting the dewpoint data below -72C has nearly no impact on the other variables.

The chilled mirror sensors are subject to flooding on rapid descents of the aircraft into the humid boundary layer. This results in temporary loss of the instruments ability to measure the dewpoint, which may last from 3 to 15 minutes, depending on conditions. This problem can also be seen in the form of "ringing", or a decaying sinusoidal oscillation of the signal, that appears after altitude changes, for example those following a period of cold soaking at high altitudes. During these periods it is advised to compare the data from both chilled mirrors and choose the one that recovers faster.

VCSEL Hygrometer was deployed for measuring atmospheric water vapor content throughout the troposphere and lower stratosphere using high sensitivity optical absorption methods, using a new, near-infrared, vertical cavity surface emitting laser (VCSEL) at 1854 nm. In conjunction with a compact, multipass, open air cell and digital signal processor (DSP) electronics, this sensor consumes very low power (< 5 W), is lightweight (< 2 kg excluding the inlet housing), and occupies only the space within an aperture plate. The use of the 1854 nm VCSEL allows for a limit of detection of <1 ppmv, a precision of 3% or 0.05 ppmv max, and a minimum sampling frequency of up to 25 Hz.

| DPLS | Dewpoint/frostpoint for left fuselage cooled-mirror sensor |
| --- | --- |
| DPLC | Dewpoint for left cooled-mirror sensor |
| DPRS | Dewpoint/frostpoint for right cooled-mirror sensor |
| DPRC | Dewpoint for right cooled-mirror sensor |
| DPXC | Reference dewpoint. In HIPPO-4 and 5 DPXC is normally set to DP_VXL. This is the value that is used for producing all humidity corrected variables. |
| MR | Mixing ratio (g/kg) based on DPXC |
| CONC_H2O_VXL | VCSEL Moisture Number Density |
| DP_VXL | Dewpoint calculated from the VCSEL H2O number densities using the improved calibration coefficients. |
| DPV_VXL | Dewpoint calculated from the VCSEL H2O mixing ratio and ATX onboard VCSEL hygrometer using fixed calibrations. This value is subject to correction when improved calibrations are produced in the lab after the project. |

RAF recommends using DPXC as a slow 'tropospheric' variable, and DP_VXL or CONC_H2O_VXL as a fast-response variable.

Attack and Sideslip: 
Measurements of attack and sideslip were done using the 5-hole nose cone pressure sensors, ADIFR and BDIFR.  Although sampled at 50 sps, internal filtering in the Mensor pressure sensors (model 6100) limits usefulness of high-rate analysis to about 5 Hz.

| ADIFR | Vertical differential pressure |
| --- | --- |
| AKRD | Attack angle. Determined from the vertical differential pressure of the radome gust probe. |
| BDIFR | Horizontal differential pressure |
| SSLIP | Sideslip angle. Determined from the horizontal differential pressure of the radome gust probe. |

Both AKRD and SSLIP were calibrated using in-flight maneuvers. However, the challenging altitude profiles of HIPPO have revealed that the parameterization of the attack calculation is imperfect and is a function of pressure altitude and airspeed that can  in turn be the function of ambient temperature. Several flights with level legs at intermediate altitudes revealed the need for further re-analyses of AKRD. The primary impact from this will be the vertical wind speed, WIC, which does not average at the expected zero during climbs and descents. Future algorithm improvements may address this issue.

Note that icing may affect the radome holes. In case the vertical differential pressure sensor holes are iced over ADIFR becomes unavailable, consequently rendering the research dependent 3D wind calculations invalid. On several HIPPO flights once the radome was affected by icing, it only recovered after another icing/thawing episode on the subsequent dip.

Dependency of the true air speed (TASX) on the radome gust probe performance through the correction to the absolute pressure measurement (PSXC) has been removed and replaced with an independent parameterization of PSX. Therefore, TASX and variables that depend on (particularly the aerosol concentrations) it will no longer be affected by the radome icing.

True airspeed: 
True airspeed was also measured primarily using a Mensor 6100 sensor, thus limiting the effective response to 5 Hz.

The radome pitot tube system uses the center hole of the 5-hole nose cone in conjunction with the research static pressure ports on the fuselage aft of the entrance door.  A standard avionics pitot tube is also mounted on the fuselage aft of the radome, and this system is also referenced to the fuselage static ports aft of the main entrance door.  It was found during empirical analysis that the fuselage pitot system gave more consistent results in reverse-heading maneuvers; it is suspected that this is due to random pressure changes at the radome center hole as has been suggested by modeling.  The fuselage system is used for the calculation of the aircraft true airspeed, as well as for attack and sideslip angles.  True airspeed is also provided from the aircraft avionics system, but this system is considered of slower response.  Measurements using the radome and fuselage pitot systems were corrected using in-flight maneuvers.

| TASR | True airspeed using the radome system |
| --- | --- |
| TASF | True airspeed from the fuselage pitot system |
| TASHC | True airspeed using the fuselage pitot system and adding humidity corrections to the calculations; this is mainly of benefit in tropical low-altitude flight |
| TAS_A | True airspeed from the avionics system |
| TASX | Reference true air speed. This is normally equal to TASF but TAS_A may be substituted in cases where TASF is compromised for any reason. This would be noted in the individual flight reports. |

RAF recommends using TASX as the aircraft true air speed.

Position and ground speed: 
The measurement of aircraft position (latitude, longitude and geometric altitude) and aircraft velocities relative to the ground are done using several sensors onboard the GV.

Novatel Omnistar-enabled GPS (Reference):  These data are sampled at 10 sps and averaged to 1 sps.  Omnistar-corrected measurements are available in real-time but accuracy may vary depending on the location. Generally, Omnistar-corrected position is accurate within 15 cm vertically and 10 cm horizontally, which was proven by comparing against a differential GPS measurements.  The values from this sensor start with a "G"; e.g.:

| GGLAT | Latitude (recommended for general use) |
| --- | --- |
| GGLON | Longitude (recommended for general use) |
| GGALT | Geometric altitude (recommended for general use) |
| GGSPD | Ground speed |
| GGVNS | Ground speed in north direction |
| GGVEW | Ground speed in east direction |
| GGQUAL | Quality factor of the Novatel GPS. Five is the highest quality with Omnistar HP correction, 6 cm horizontal, 12 cm vertical specified accuracy. Two is the lower quality Omnistar VBS corrected data with a sub 1 m specified accuracy, this level is usually seen on the edge of the Omnistar HP coverage area; One is the lowest quality, GPS-only position with a 10-15 m specified accuracy. Nine is a WAAS-augmented GPS with a 8 m accuracy. All estimates are based on a 99% confidence limit. |
| ALTX | Reference altitude. This value is recommended for use for most applications. It is set to be equal to GGALT for most flights. In rare cases where GGALT data were interrupted, the value will be substituted by the best available alternative such as GGALT_GMN (see explanation below). |

These are good values to use for cases where the highest real-time accuracy is needed.  These variables are subsequently used to constrain the INS drift for the calculations of the GV winds; more about this below.

A secondary Garmin GPS system provided redundant position measurements that should be used during periods of noisy or missing Novatel Omnistar-corrected GPS data. The Garmin data are not corrected and the accuracy is within 10-15 meters. These variables have the same naming convention as the reference GPS above but are distinguished by a suffix "_GMN".

Honeywell inertial reference system 1 and 2:  The GV is equipped with three inertial systems.  Data from the first two of these are logged on the main aircraft data logger, with subscripts the latter having variable names with suffix "_IRS2".  The advantage of the IRS values is that they typically have very high sample rates and very little noise from measurement to measurement.  However, since they are based on accelerometers and gyroscopes, their values may drift with time.  The drift is corrected for by filtering the INS positions towards the GPS positions with a long time-constant filter; the filtered values have a "C" added to the end.

| LAT | latitude from IRS 1, no GPS filtering |
| --- | --- |
| LATC | latitude from IRS 1, filtered towards GPS values |
| LAT_IRS2 | latitude from IRS 2, no GPS filtering |
| LON | longitude from IRS 1, no GPS filtering |
| LONC | longitude from IRS 1, filtered towards GPS values |
| LON_IRS2 | longitude from IRS 2, no GPS filtering |
| GSF | ground speed from IRS 1, no GPS filtering |
| GSF_IRS2 | ground speed from IRS 2, no GPS filtering |

The choice of variables for position analysis depends on the type of analysis; in general the Novatel Omnistart GPS is the most accurate and preferred. RAF recommends using ALTX, GGLAT and GGLON for position information for most of the flights, with exceptions noted specifically in the individual flight comments below.

Please note that when the GPS position is lost, IRS data are also lost. When the system later recovers, LATC, LONC and ALTC show deviations at the edges of the data gap. This is an artifact of the filtering algorithm and should be ignored.

Not all INS variables are output in the final data set, including IRS2. If you require more detailed INS data please contact RAF.

  Attitude angles: 
Aircraft attitude angles are measured by the two Honeywell IRS units.

| PITCH | Pitch of the aircraft |
| --- | --- |
| PITCH_IRS2 | Pitch of the aircraft from the second IRS |
| ROLL | Roll of the aircraft |
| ROLL_IRS2 | Roll of the aircraft from the second IRS |
| THDG | True heading of the aircraft |
| THDG_IRS2 | True heading of the aircraft from the second IRS |

The values of pitch angle (PITCH) have been corrected using in-flight measurements to give approximately the same values as the aircraft attack angle (AKRD) for long, straight and level parts of flights; this correction is performed to give a near-zero mean updraft (WIC) over extended flight legs.  The variation from flight to flight of this offset is caused by small differences in the pre-flight alignment of the inertial navigation system.  No alignment correction has been applied to PITCH_IRS2.

  Wind speeds: 
Wind speeds are derived from the 5-hole radome gust probe combined with other pressure, temperature and inertial measurements supported by GPS data. The Mensor 6100 pressure sensor (ADIFR, BDIFR and radome dynamic pressure QCF) provides pressure measurements sampled at 50 sps and thus resulting in power spectra to 25 Hz.  Examination of power spectra and specifications from Mensor indicate that the sensors have internal filters with a -3dB (half-power) cutoff at 12 Hz, resulting in a noticeable roll-off in the spectra beginning approximately at 6 to 7 Hz.  Users of wind data should be aware that contributions to covariances and dissipation calculations will be affected at and above these frequencies.

The following lists the most commonly used wind variables:

| UI | Wind vector, east component |
| --- | --- |
| UIC | Wind vector, east component, GPS corrected for INS drift |
| VI | Wind vector, north component |
| VIC | Wind vector, north component, GPS corrected |
| UX | Wind vector, longitudinal component |
| UXC | Wind vector, longitudinal component, GPS corrected |
| VY | Wind vector, lateral component |
| VYC | Wind vector, lateral component, GPS corrected |
| WI | Wind vector, vertical gust component |
| WIC | Wind vector, vertical gust component, GPS corrected |
| WS | Wind speed, horizontal component |
| WSC | Wind speed, horizontal component, GPS corrected |
| WD | Horizontal wind direction |
| WDC | Horizontal wind direction, GPS corrected |

WIC is a derived parameter calculated from the attack angle, which is calculated from the lateral and vertical pressure measurements from the radome 5-hole gust sensor. It was understood in HIPPO that the attack calculation requires refinement. It is expected that AKRD and WIC calculation will change in the future releases of the data to better reflect the vertical wind conditions, resulting in the average zero vertical wind speed during the altitude changes.

RAF recommends using the GPS corrected wind components, i.e., the variables ending in "C". Uncorrected wind variables may not be included in the final data release.

Liquid water content: 
A PMS-King type liquid water content sensor was installed on the GV. The corrected liquid water content obtained from the power consumption required to maintain a constant temperature, taking into account the effect of convective heat losses.

| PLWCC | Liquid water content derived from PLWC, g/m3 |
| --- | --- |
| PLWC | Raw dissipated power, watt |

Note that the King probe is not optimized for the GV airspeeds. Significant improvements have been made to the processing code to better quantify changes in the heat transfer related to the changes in the airspeed. However, rapid fluctuations in the PLWCC baseline are still observed in the PLWCC that are not realistic. Please compare the PLWCC data with PLWCD_* that is calculated from the cloud droplet probe to assess the quality of the liquid water data from the King probe.

Icing rate indicator

Rosemount Model 871FA Icing Rate Detector was installed on the right outboard wing pylon access plate (RICE). The instrument reports a sharp spike when icing is detected, often coincident with a diversion between QCR and QCF, indicative of the radome tip icing.

The Microwave Temperature Profiler (MTP) was deployed. The MTP is a passive microwave radiometer, which measures the natural thermal emission from oxygen molecules in the earth’s atmosphere for a selection of elevation angles between zenith and nadir by scanning through an arc in the flight direction. The MTP observing frequencies are located on oxygen absorption lines at 56.363, 57.612 and 58.363 GHz. The measured brightness temperatures versus elevation angle are converted to air temperature versus altitude using a quasi-Bayesian statistical retrieval procedure. An altitude temperature profile (ATP) is produced in this manner every 18 seconds along the flight track. Temperature accuracy is approximately 1 K within 3 km of flight altitude, and < 2K within 6 km of flight altitude. Vertical resolution is approximately 150 meters at flight level, and approximately half the distance from the aircraft away from flight level. Data are provided in the form of altitude temperature profiles and color-coded temperature “curtains”, and may be converted to isentropic cross sections.

CDP (Cloud Droplet Probe)

The CDP is a commercial instrument from Droplet Measuring Technologies (DMT). It measures the intensity of forward light scattering (4 – 12°) to determine the sizes of individual cloud droplets. An internal multi-channel analyzer assigns individual particles to bins, and the data interface outputs a histogram of particle size and concentration. On the NSF/NCAR C-130 and HIAPER, the CDP is mounted in a PMS canister. Variables output by the CDP instrument have _LWI suffix.

For HIPPO-4 and 5, the CDP real time size distributions and early post processing results have been affected by an incorrect binning. This issue has been corrected and will have caused a slight shift of size distributions and derived liquid water contents; the correct LWC from CDP is 10-15% lower than originally calculated.

| size range | 2 – 50 µm diameter |
| --- | --- |
| concentration range | 0 – 5,000 cm-3 |
| number of size bins | 10, 20, 30, or 40 |
| sample area | 200µm x 1.5mm |
| volume sample rate | 30 cm3/s at airspeed 100 m/s |
| airspeed range | 10 – 200 m/s |
| data interface | serial RS-232 or RS-422 |
| data rate | 10 histograms per second |

more information: http://www.dropletmeasurement.com/products/CDP.htm

# Ultra-High Sensitivity Aerosol Spectrometer – UHSAS

The UHSAS is a single-particle light scattering instrument. It uses a CW high energy laser diode, wide angle collection optics centered at 90°, and four stages of amplification to size aerosol particles according to their scattered light. It assigns bins to individual particles and outputs a histogram of particle size and concentration. The RAF version of this instrument has been highly modified from the commercially-available lab bench version.  It uses volume flow controllers to keep the flow constant over a wide range of operating pressures and temperatures. The probe mounts in a PMS canister. The variables output by UHSAS have the suffix _RWO. A Labview control program running on a cabin-mounted laptop computer is used to calibrate and properly initialize the instrument in preflight or during the flight as necessary.

| size range | 75 – 1000 nm diameter |
| --- | --- |
| concentration range | 0 – 18,000 cm-3 |
| number of size bins | 99 |
| volume sample rate | 1 cm3/s |
| data interface | serial RS-232 |
| data rate | 10 histograms per second |

more information: http://www.dropletmeasurement.com/products/UHSAS.htm

Data logging and averaging:

Analog data were logged at 500 sps and averaged to 1 sps. Serial data (e.g., RS-232), ARINC data (IRS units), etc. were recorded at the instrument-specific output rate.

The recordings listed for a given second contains measurements logged at e.g., 12:00:00.000 and until 12:00:01.  The value of "Time" corresponding to this interval is given a 12:00:00 in the released data set.

All measurements are "time-tagged" at the time of logging.  Subsequently these measurements are interpolated onto a regular grid and averaged.

RAF staff have reviewed the data set for instrumentation problems.  When an instrument has been found to be malfunctioning, specific time intervals are noted.  In those instances the bad data intervals have been filled in the netCDF data files with the missing data code of -32767.  In some cases a system may be out for an entire flight.

# Calibrations

| Sensor | Type | DSM | S/N | Polynomial coefficients |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | A | B*x | C*x2 | D*x3 |
| PSF | Absolute | 304 | 92028 | -0.30806 | 1.0002 | 7.335E-8 |  |
| TTFH1 | Rosemount | 305 | A50738 |  |  |  |  |
| TTFH2 | Rosemount | 305 | A50738 |  |  |  |  |
| TTHR1 | Harco | 305 | 708094 |  |  |  |  |
| TTHR2 | Harco | 305 | 708094 |  |  |  |  |
| QCR | Mensor | 305 | 590684 | -0.1894 | 1.0006 | 1.266E-6 |  |
| QCF | Mensor | 305 | 590682 | 0.075274 | 1.0003 | 1.905E-6 |  |
| ADIFR | Mensor | 305 | 590688 |  |  |  |  |
| BDIFR | Mensor | 305 | 590686 | -0.00204 | 0.99689 | 2.337E-7 |  |
| DPLS | Buck | 1011C |  |  |  |  |  |
| DLRS | Buck | 1011C |  |  |  |  |  |

The following table identifies the sensors serial numbers and calibrations that were used on the GV for HIPPO Global Phase 4:

The following table identifies the sensors serial numbers and calibrations that were used on the GV for HIPPO Global Phase 5:

| Sensor | Type | DSM | S/N | Polynomial coefficients |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | A | B*x | C*x2 | D*x3 |
| PSF | Absolute | 304 | 92028 | -0.30806 | 1.0002 | 7.335E-8 |  |
| TTFH1 | Rosemount | 305 | A50738 |  |  |  |  |
| TTFH2 | Rosemount | 305 | A50738 |  |  |  |  |
| TTHR1 | Harco | 305 | 630393 |  |  |  |  |
| TTHR2 | Harco | 305 | 630393 |  |  |  |  |
| QCR | Mensor | 305 | 590684 | -0.085285 | 1.0007 | -2.849E-6 |  |
| QCF | Mensor | 305 | 590682 | -0.015201 | 1.0003 | 1.034E-6 |  |
| ADIFR | Mensor | 305 | 590688 |  |  |  |  |
| BDIFR | Mensor | 305 | 590686 | -0.029908 | 0.99725 | 2.941E-6 |  |
| DPLS | Buck | 1011C |  |  |  |  |  |
| DLRS | Buck | 1011C |  |  |  |  |  |

Calibration data are not provided for temperatures because of the continuing work on the absolute calibrations that affects in flight calibration coefficients.

Data Quality Control

General:

- Vertical wind speed (WIC) data are reliable only during straight and level flight. Expect deviations whenever the aircraft is not in straight and level flight. For HIPPO there was very little straight and level flight and effort was made to tune the WIC processing to obtain good vertical wind in climbs and descents but deviations can still be seen.

- DPRC, DPLC: dewpointers tend to overshoot and oscillate after rapid temperature increase on aircraft descents. Using best judgment, these overshoots are removed from DP..C variables, but are left in DP..S, which are the source variables for DP..C, for comparison. The operating range for both DPLC and DPRC is down to approximately -70C and values below that should not be used for quantitative analyses. DP_VXL (dewpoint measurement produced by VCSEL hygrometer using post calibrations) is not affected by rapid  temperature changes and is normally used as DPXC.

- Icing was an infrequent occurrence during HIPPO-4. To detect icing conditions look for the onset of a difference between QCR and QCF in addition to spiking signals in PLWCC and RICE. The radome tip hole (QCR) ices over very quickly in icing conditions, resulting in the lag in response or constant pressure reading. These data are not representative of flight conditions (refer to QCF instead) but are useful for identifying periods of icing early, so they were left in the dataset. Variables dependent on the QCR (for example TASR) will also be unusable during radome icing.

- Sometimes the icing has been extensive enough to reach the vertical differential pressure sensor holes on the radome (ADIFR). This measurement is critical for determining the air flow angles for wind calculations and when the radome holes freeze over the wind data become unavailable.

- PLWCC processing: the background output of the King hot wire probe is loosely coupled to zero using a sliding window average. Once a liquid water cloud is encountered the coupling unlatches and the liquid water content is calculated using the last background value as a reference. This sometimes results in artifacts if the baseline continues to drift as a function of airspeed during the cloud encounter since the baseline at the exit time will try to return to the last known good value. This is an algorithm feature that will be addressed in the future through refinements in the PLWCC airspeed parameterization.

- 2D-C probe located on LWO has demonstrated elongated cloud droplet shapes once the droplets reach about 1 mm in diameter. The cause for this may be related to the air compression and flow artifacts near the leading edge of the wing and tip of the PMS can. This requires further investigation and more data from clouds. Please keep in mind that elliptical shapes of the particles can lead to overestimation of the liquid water content derived from the 2D probe data.

- UHSAS probe, while evidently operating fine with sample flow well controlled, may intermittently report erroneous sample flow. Since the sample flow is directly used for concentration and size distribution calculations, this resulted in stairstep changes in these data. Thorough analyses of the UHSHAS data revealed that raw counts did not change during apparent flow changes, and none of the other housekeeping parameters were affected. RAF concluded that the flow was in fact correct and the readout of the flow was misreported by the probe. This allowed the average sample flow, which was 0.728 sccs with Ϭ=0.00028, to be used in places where the flow was reported incorrectly.

Flight specific notes, HIPPO-4:

RF01: No specific issues.

RF02: No specific issues.

RF03: Radome icing was detected from 22:23 to 22:32, resulting in the loss of QCR data.

RF04: Novatel GPS data gap from 07:31:41 to 07:32:45. The data from the redundant Garming GPS (GG*_GMN) are available during this gap. ALTX is set to the Garmin GPS position data. Please note that the Garmin GPS is not Omnistar-capable and the accuracy of the position is ±15m as opposed to ±0.5m for the primary GPS. If you need high accuracy please use GGALT, GGLAT and GGLON for most of the flight and only substitute GGALT_GMN, GGLON_GMN and GGLAT_GMN during the gap.

RF05: Novatel GPS data gap from 23:11:26 to 23:17:25. The data from the redundant Garming GPS (GG*_GMN) are available during this gap. ALTX is set to the Garmin GPS position data. Please note that the Garmin GPS is not Omnistar-capable and the accuracy of the position is ±15m as opposed to ±0.5m for the primary GPS. If you need high accuracy please use GGALT, GGLAT and GGLON for the entire flight and only substitute GGALT_GMN, GGLON_GMN and GGLAT_GMN during the gap.

RF06: Radome icing was detected from 05:39:36 to 06:12:09. QCR data are unusable during this period, wind measurements are not affected.

RF07: No specific issues.

RF08: No specific issues.

RF09: Radome icing was detected from 15:40:31 to 15:45:32 and from 06:25:49 to 06:31:54. QCR data are unusable during this period, wind measurements are not affected.

RF10: Radome icing was detected from 16:22 to 16:43 and from 20:03 to 21:02. QCR data are unusable during this period, wind measurements are not affected.

RF11: No specific issues.

RF12: Obstruction by water or ice of the radome gust probe lines caused ADIFR to lose signal from 23:41:16 to 24:35:59 resulting in the loss of derived 3D wind data.

Flight specific notes, HIPPO-5:

RF01: Right chilled mirror dewpointer (DP_DPR, DPRC) was inoperable for the entire flight.

RF02: Short ADIFR and QCF outage from 21:12:16 to 21:13:32 resulted in a data gap in derived wind speed, TASX and aerosol concentration variables. Due to a very short gap duration TASX was not replaced with TAS_A for this flight; if this becomes an issue the flight can be reprocessed using TAS_A that will allow recovery of the aerosol concentration data for the affected period.

RF03: ADIFR and QCF have numerous small data gaps throughout the entire flight. TASX is replaced by TAS_A for this flight to recover missing aerosol concentration data for these gaps. Wind data have the same gaps and not recoverable because the attack angle (AKRD) is directly calculated from the affected ADIFR. UHSAS concentration data are also not available for these small gaps because the UHSAS flow was being recorded on the same DSM that was affected.

RF04: No specific issues.

RF05: Radome icing was encountered from 02:19 to 02:29 resulting in the loss of QCR data. This has not impacted any other measurements but please take a note of the icing event in case MTP or other optical probes were impacted.

RF06: Data gap in ADIFR from 18:25:21 to 19:31:16 resulted in the loss of dependent wind variables. Suspected cause: water ingestion on climb-out and subsequent freezing in the transducer tubing.

RF07: The primary GPS lost signal for parts of the flight. The reference altitude (ALTX) is set to the secondary GPS. Please note that this GPS is less accurate than the primary (15 m as opposed to 50 cm). If high accuracy is required please use GGALT, GGLAT and GGLON for most of the flight and only substitute *_GMN variables during the periods of missing primary GPS data. UHSAS did not operate properly on this flight; size distributions only show particles in the smallest bins and are not representative of real conditions.

RF08: The primary GPS lost signal for small part of the flight. The reference altitude (ALTX) is set to the secondary GPS. Please note that this GPS is less accurate than the primary (15 m as opposed to 50 cm). If high accuracy is required please use GGALT, GGLAT and GGLON for most of the flight and only substitute *_GMN variables during the periods of missing primary GPS data. UHSAS did not operate properly on this flight; size distributions only show particles in the smallest bins and are not representative of real conditions.

RF09: Cloud droplet probe (CDP) was inoperable for the entire flight.

RF10: ADIFR sensor blockage caused a data gap from 22:09:45 to 23:03:07. QCR data loss indicates icing of the radome from 00:08:43 to 00:36:32 and 01:39:49 to 01:57:34

RF11: Radome icing was detected from 21:34:53 to 21:41:01 and from 02:52:29 to 03:09:24. QCR data are lost; other variables are not affected.

RF12: Left chilled mirror dewpointer (DP_DPL, DPLC) was inoperable for the entire flight. Radome icing was encountered from 19:47:06 to 20:37:10 and from 21:34:28 to the end of flight. QCR was affected; ADIFR and the derived wind data were not.

RF13: Radome icing was encountered from 19:47:06 to 20:37:10 and from 21:34:28 to 25:35:00. QCR data lost; no other data affected.

FF01: No specific issues.